Run plan: EJ312-P Cup PMT bias linearity test (2019-03-18)

In order to find a suitable operating bias for the EJ312P cup PMT, experiments are needed. Using a source of single species particles (gammas or neutrons) the bias will be systematically changed and the counts above threshold recorded. The goal is to find a plateau of “stability” where the counts are stable across a range of bias.
A gamma source (Cs-137, Co-60 or Na-22) can be used for the gamma tests.
The AmBe (or Pube) source can be used for the neutrons tests. However lead shielding is needed to block most of the gamma-rays which will ensure an almost “clean” spectrum of neutrons

In order to study whether or not the Compton edge can be improved in terms of resolution, two detector can be used in logic AND and used to trigger the data acquisition of the first detector as a function of scattering angle of the gamma ray.

Material:
- Gamma source (Cs-137, Co-60 or Na-22)
- Neutron source (AmBe or PuBe) Cf?
- NE213 cup detector
- NIM crate
- Discriminator (CFD) module (for trigger lvl)
- Visual scaler or MCA (to get number of events)

Procedure:
1) Determine CFD threshold (ideas where to put it?)
2) Gamma source: Run 5-10 min, vary bias from -1400 V to -2300 V in steps of 50 V. Note down voltage and current draw.
3) Neutron source: Run 5-10 min, vary bias from -1400 V to -2300 V in steps of 50 V. Note down voltage and current draw.

Results:
- To produce count rate vs bias plots and determine a good operating bias for the PMT.

Estimated time:
Setup: 1h
Test/verification/modification: 1h
Run: Needs “babysitting”. Prob 2 days for all measurements.